**ELK** **NERR Water Quality Metadata**

**January to December 2017**

**Latest Update:** January 22, 2019

**I. Data Set and Research Descriptors**

**1) Principal investigator(s) and contact persons**

Kerstin Wasson, Research Coordinator Elkhorn Slough NERR

[kerstin.wasson@gmail.com](mailto:kerstin.wasson@gmail.com) 1700 Elkhorn Rd

John Haskins, Water Quality Scientist Watsonville, CA 95076

[john@elkhornslough.org](mailto:john@elkhornslough.org) (831) 728-2822

Rikke Jeppesen, Estuarine Ecologist

[rikke@elkhornslough.org](mailto:rikke@elkhornslough.org)

**2) Entry verification**

Deployment data are uploaded from the YSI data logger to a Personal Computer (IBM compatible). Files are exported from EcoWatch in a comma-delimited format (.CDF) , EcoWatch Lite in a comma separated file (CSV) or KOR Software in an Excel File (.XLS) and uploaded to the CDMO where they undergo automated primary QAQC; automated depth/level corrections for changes in barometric pressure (cLevelor cLevel parameters); and become part of the CDMO’s online provisional database. All pre- and post-deployment data are removed from the file prior to upload. During primary QAQC, data are flagged if they are missing or out of sensor range. The edited file is then returned to the Reserve for secondary QAQC where it is opened in Microsoft Excel and processed using the CDMO’s NERRQAQC Excel macro. The macro inserts station codes, creates metadata worksheets for flagged data and summary statistics, and graphs the data for review. It allows the user to apply QAQC flags and codes to the data, remove any overlapping deployment data, append files, and export the resulting data file for upload to the CDMO. Upload after secondary QAQC results in ingestion into the database as provisional plus data, recalculation of cLevelor cLevel parameters, and finally tertiary QAQC by the CDMO and assimilation into the CDMO’s authoritative online datSabase. Where deployment overlap occurs between files, the data produced by the newly calibrated sonde is generally accepted as being the most accurate. For more information on QAQC flags and codes, see Sections 11 and 12.

**3) Research objectives**

The goal of the research and monitoring of water quality at Elkhorn Slough NERR is to establish baselines for water quality parameters for Elkhorn Slough by using South Marsh as a control site while monitoring two impacted sites (Azevedo Pond and North Marsh) for possible problems. Additionally, in order to identify oceanic influence on the water quality parameters at Elkhorn Slough, we monitor a fourth site, Vierra Mouth (VM). Water quality measurements are recorded every 15 minutes over a four week period at the four sites in Elkhorn Slough. One site, South Marsh (SM), is in a relatively un-impacted side channel of the slough and the second site, Azevedo Pond (AP), is in a pond that receives fertilizer and pesticide run-off from an adjoining strawberry field. The third site, North Marsh (NM), was added in April 1999 and is located in an area where there is both agricultural and non-agricultural run-off. The fourth site, Vierra Mouth, is located at the mouth of the slough and is used to identify oceanic influence. This site was added March 14, 2001.

**4) Research methods**

The Elkhorn Slough water quality monitoring program began in July 1995. All four sites described above are monitored simultaneously. Prior to YSI deployment at SM, a 20-foot length of 4-inch diameter PVC pipe was placed in the slough to house the YSI. Holes were drilled in the pipe to remove 10% of the pipe's surface area and to allow water flow across the YSI. The pipe was positioned vertically in the slough with one end pushed into the soft bottom sediments and the other end secured to a permanent dock. A bolt inside and across the diameter of the pipe maintains the YSI about one foot (30cm) above the bottom. The setup is exactly the same at North Marsh except that the PVC pipe is secured to shore. Prior to YSI deployment at AP, a supportive framework of rope and PVC was placed at the site to house the YSI. The structure maintained the YSI exactly one foot (30cm) above the pond bottom. In 2012, a 40 ft board walk and a 5 ft black ABS pipe were installed to replace the original sonde support structure. A larger diameter ABS pipe slides up and down the 2-inch ABS pipe, which is protruding vertically from the substrate. Perpendicularly mounted on the sliding ABS is a 4-inch diameter white PVC pipe, 3-ft long, which functions as the housing for the YSI. The white PVC housing has been drilled with 1-inch diameter holes, in order to ensure water circulation at the YSI. At Vierra Mouth, the installed PVC pipe housing the YSI sonde is the same as at South Marsh. Every 15 minutes over a 30-day period, measurements of specific conductivity, salinity, dissolved oxygen, temperature, depth, turbidity, and pH are recorded. At South Marsh, additionally chlorophyll is recorded.

At the end of each approximate 28-day period, the YSI datalogger is brought back to the lab to download the data, and to clean and recalibrate the sonde. Data are downloaded onto a PC and then all data are transferred to a server at the reserve ([\\SBSERVER](file:///\\SBSERVER)). Sonde body and probes are cleaned. Calibration is performed as outlined in the YSI manual. Buffer solutions of pH 7 and pH 10 are purchased from a scientific supply store and used for two-point calibration of the pH probe. Since the beginning of July 1999 a conductivity standard of 53 mS/cm was used to calibrate the conductivity probe. The turbidity probe is calibrated using a two-point calibration with DI water (0 NTU) and 124 NTU standard. Deionized water is used to calibrate the chlorophyll probe on the sonde at SM. The DO probe was calibrated using a 2-point calibration. A mixture of 2 g sodium sulfite dissolved in 1 L tap water was left for 1 hour to equilibrate at 0% saturation. Tap water saturated with an air-stone for at least 15 minutes prior to calibration was used for the 100% saturation. Before retrieving a sonde from a site a new sonde is calibrated in the lab before retrieving the currently deployed one, in order to replace the one in the field. This eliminates the loss of data due to cleaning and calibration. QA/QC was done according to the CDMO manual by John Haskins and/or Rikke Jeppesen using the macros provided by the CDMO. Data are then looked over more rigorously to identify and document anomalies and missing data. Additionally, John Haskins uses Matlab, in order to better identify anomalies in the monthly data, by overlaying the current month’s data onto all data at a site, since 1995. This method allows to more easily identify probe drift and malfunction, in addition to natural variation in the data. This method has only been used for all four sites since 2011.

A Sutron Sat-Link2 transmitter was installed at the Azevedo Pond station on 04/22/2014 and transmits data to the NOAA GOES satellite, NESDIS ID #3B053520. 3B053520 is the GOES ID for that particular station. The transmissions are scheduled hourly and contain four (4) data sets reflecting fifteen minute data sampling intervals. Upon receipt by the CDMO, the data undergoes the same automated primary QAQC process detailed in Section 2 above. The “real-time” telemetry data become part of the provisional dataset until undergoing secondary and tertiary QAQC and assimilation in the CDMO’s authoritative online database. Provisional and authoritative data are available at [http://cdmo.baruch.sc.edu](http://cdmo.baruch.sc.edu/).

A Sutron Sat-Link2 transmitter was installed at the South Marsh station on 09/28/2006 and transmits data to the NOAA GOES satellite, NESDIS ID #3B026768. 3B026768 is the GOES ID for that particular station. The transmissions are scheduled hourly and contain four (4) data sets reflecting fifteen minute data sampling intervals. Upon receipt by the CDMO, the data undergoes the same automated primary QAQC process detailed in Section 2 above. The “real-time” telemetry data become part of the provisional dataset until undergoing secondary and tertiary QAQC and assimilation in the CDMO’s authoritative online database. Provisional and authoritative data are available at [http://cdmo.baruch.sc.edu](http://cdmo.baruch.sc.edu/).

A Sutron Sat-Link2 transmitter was installed at the North Marsh station on 07/17/2017 and transmits data to the NOAA GOES satellite, NESDIS ID #3B0543B0 is the GOES ID for that particular station. The transmissions are scheduled hourly and contain four (4) data sets reflecting fifteen minute data sampling intervals. Upon receipt by the CDMO, the data undergoes the same automated primary QAQC process detailed in Section 2 above. The “real-time” telemetry data become part of the provisional dataset until undergoing secondary and tertiary QAQC and assimilation in the CDMO’s authoritative online database. Provisional and authoritative data are available at [http://cdmo.baruch.sc.edu](http://cdmo.baruch.sc.edu/).

A Sutron Sat-Link2 transmitter was installed at the Vierra Mouth station on 6/27/2012 and transmits data to the NOAA GOES satellite, NESDIS ID #3B04D428. 3B04D428 is the GOES ID for that particular station. The transmissions are scheduled hourly and contain four (4) data sets reflecting fifteen minute data sampling intervals. Upon receipt by the CDMO, the data undergoes the same automated primary QAQC process detailed in Section 2 above. The “real-time” telemetry data become part of the provisional dataset until undergoing secondary and tertiary QAQC and assimilation in the CDMO’s authoritative online database. Provisional and authoritative data are available at [http://cdmo.baruch.sc.edu](http://cdmo.baruch.sc.edu/).

**5) Site location and character**

Azevedo Pond (AP)(36°50’44.64”N, 121°45’13.24”W) is in a pond that receives fertilizer and pesticide run-off from a strawberry field in year-round production. The sample station is located about 10m from a tidal control structure in front of a culvert connecting the pond to the slough. In 2017, the tide ranged from 1.18 to 2.48 meters at this site and salinity ranged from 0.8 ppt during heavy run-off to 41.0 ppt during strong evaporation. The YSI sonde associated with this site (collecting readings for the water quality dataset) is located approximately 30 cm off the bottom, which is composed of silty mud. An EXO2 sonde is deployed at this site.

North Marsh (NM)(36°50’04.75”N, 121°44’18. 33”W) is located in-between South Marsh and Azevedo Pond. This site is impacted by both agricultural and urban run-off. In 2017, the tide ranged from approximately 0.54 to 1.23 meters at this site. Salinity ranged between 7.8 and 46.8 ppt and is affected by freshwater run-off from agriculture and upland run-off. The YSI sonde associated with this site (WQ dataset) is approximately 30 cm off the bottom, which is composed of silty mud. An EXO2 sonde is deployed at this site.

South Marsh (SM)(36°49’05.00”N, 121°44’21.83”W) which is located approximately 3 km south of NM and is surrounded by mostly reserve land, is in a side channel of the slough and is relatively free from impact by anthropogenic influence. This site receives run-off mostly from uplands with some run-off coming from cattle ranches. This site receives the least amount of pollution. In 2017, the tidal range was from -0.52 to 2.82 meters at this site and the salinity range was from 10.6 to 35.7 ppt. The YSI sonde associated with this site (collecting readings for the water quality dataset) is approximately 30 cm off the bottom, which is composed of compacted silty mud.

An EXO2 sonde is deployed at this site.

The fourth site Vierra Mouth (VM) (36°48’39.95”N, 121°46’45.22”W) is located at the mouth of the slough and is used to identify oceanic influence. In 2017, the tidal range was from -0.26 to 2.39 meters at this site and salinity ranged from 12.2 to 34.6 ppt. The YSI sonde associated with this site (collecting readings for the water quality dataset) is located approximately 30 cm off the bottom which is composed of compacted mud and sand due to strong tidal currents. This site receives drainage from the entire watershed due to its location at the mouth. There are several auto wrecking yards located approximately 2 km east of this site. An EXO2 sonde is deployed at this site.

SWMP Station Timeline

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Station Code | SWMP Status | Station Name | Location | Active Dates | Reason Decommissioned | Notes |
| AP | P | Azevedo Pond | 36°50’44.64”N, 121°45’13.24”W | June 1995-current | NA | NA |
| NM | P | North Marsh | 36°50’04.75”N, 121°44’18. 33”W | April 1999 - curernt | NA | NA |
| SM | P | South Marsh | 36°49’05.00”N, 121°44’21.83”W | June 1995-current | NA | NA |
| VM | P | Vierra Mouth | 36°48’39.95”N, 121°46’45.22”W | March 2001 - current | NA | NA |

**6) Data collection period**

Sampling at Azevedo Pond and South Marsh began simultaneously at the end of June 1995 and data collection is ongoing. North Marsh began sampling on April 15, 1999. Vierra Mouth began sampling on March 14, 2001 at 13:00.

# Deployment recovery times and dates

**Azevedo Pond**

Start Date Start Time End Date End Time

12/05/2016 08:15 01/03/2017 11:15

01/03/2017 11:30 01/30/2017 11:00

01/30/2017 11:15 02/27/2017 10:00

02/27/2017 10:15 03/27/2017 13:00

03/27/2017 13:15 04/24/2017 10:15

04/24/2017 10:30 05/22/2017 10:30

05/22/2017 10:45 06/19/2017 11:45

06/19/2017 12:00 07/17/2017 10:15

07/17/2017 10:30 08/14/2017 10:30

08/14/2017 10:45 09/11/2017 10:00

09/11/2017 10:15 10/10/2017 10:30

10/10/2017 10:45 11/06/2017 11:30

11/06/2017 11:45 12/04/2017 11:15

12/04/2017 11:30 01/02/2018 11:30

**North Marsh**

Start Date Start Time End Date End Time

12/12/2016 11:45 01/09/2017 10:30

01/09/2017 10:45 02/06/2017 11:00

02/06/2017 11:15 03/06/2017 11:00

03/06/2017 11:15 04/03/2017 11:30

04/03/2017 13:00 05/01/2017 10:00

05/01/2017 10:15 05/26/2017 11:15

05/26/2017 11:30 06/26/2017 10:00

06/26/2017 10:15 07/24/2017 08:15

07/24/2017 08:30 08/21/2017 11:15

08/21/2017 11:30 09/18/2017 09:45

09/18/2017 10:00 10/16/2017 10:30

10/16/2017 10:45 11/13/2017 11:30

11/13/2017 11:45 12/08/2017 12:15

12/08/2017 12:30 01/08/2018 11:30

**South Marsh**

Start Date Start Time End Date End Time

12/21/2016 09:15 01/25/2017 11:30

01/25/2017 11:45 02/23/2017 12:30

02/23/2017 12:45 03/20/2017 10:15

03/20/2017 10:30 04/17/2017 10:00

04/17/2017 10:15 05/15/2017 09:00

05/15/2017 09:15 06/12/2017 09:30

06/12/2017 09:45 07/10/2017 09:45

07/10/2017 10:00 08/07/2017 11:00

08/07/2017 11:15 09/05/2017 12:00

09/05/2017 12:30 10/02/2017 10:15

10/02/2017 10:30 10/30/2017 09:15

10/30/2017 09:30 11/27/2017 13:00

11/27/2017 13:15 12/20/2017 09:00

12/20/2017 09:45 01/22/2018 11:45

**Vierra Mouth**

Start Date Start time End Date End Time

12/19/2016 11:30 01/17/2017 13:00

01/17/2017 13:15 02/13/2017 11:30

02/13/2017 11:45 03/09/2017 16:45

03/09/2017 17:00 04/10/2017 08:30

04/10/2017 08:45 04/12/2017 12:30

04/12/2017 13:00 05/08/2017 09:45

05/08/2017 10:00 06/05/2017 10:30

06/05/2017 10:45 06/29 /2017 10:45

06/29 /2017 11:00 07/31/2017 09:30

07/31/2017 10:15 08/28/2017 10:15

08/28/2017 10:30 09/25/2017 11:30

09/25/2017 11:45 10/23/2017 10:00

10/23/2017 10:15 11/20/2017 10:00

11/20/2017 12:00 12/18/2017 10:45

12/18/2017 11:30 01/15/2018 10:30

**7) Distribution**

NOAA retains the right to analyze, synthesize and publish summaries of the NERRS System-wide Monitoring Program data.  The NERRS retains the right to be fully credited for having collected and process the data.  Following academic courtesy standards, the NERR site where the data were collected should be contacted and fully acknowledged in any subsequent publications in which any part of the data are used.  The data set enclosed within this package/transmission is only as good as the quality assurance and quality control procedures outlined by the enclosed metadata reporting statement.  The user bears all responsibility for its subsequent use/misuse in any further analyses or comparisons.  The Federal government does not assume liability to the Recipient or third persons, nor will the Federal government reimburse or indemnify the Recipient for its liability due to any losses resulting in any way from the use of this data.

Requested citation format:

NOAA National Estuarine Research Reserve System (NERRS). System-wide Monitoring Program. Data accessed from the NOAA NERRS Centralized Data Management Office website: <http://www.nerrsdata.org/>; *accessed* 12 October 2012..

NERR water quality data and metadata can be obtained from the Research Coordinator at the individual NERR site (please see Principal Investigators and Contact Persons), from the Data Manager at the Centralized Data Management Office (please see personnel directory under the general information link on the CDMO home page) and online at the CDMO home page [www.nerrsdata.org](http://www.nerrsdata.org).  Data are available in comma delimited format.

**8) Associated researchers and projects**

In addition to our reserve staff, multiple researchers from multiple institutions conduct research that makes use of our long-term monitoring data. The researchers are listed below.

As part of the SWMP long-term monitoring program, ES NERR also monitors 15-minute meteorological along with monthly grab samples and diel sampling for nutrient data which may be correlated with this water quality dataset. These data are available at [www.nerrsdata.org](http://www.nerrsdata.org).

**Research conducted by reserve staff**

Susie Fork conducts field survey monitoring of shorebirds; egret and heron rookery, bird nest boxes, raptors, and invertebrate populations.

Rikke Jeppesen and Susie Fork conduct annual crab trapping in order to track crab populations, particularly invasions by non-native European crabs. Additionally, Susie Fork conducts annual invertebrate surveys on mudflats, in permanent transects. The surveys include various clam and shrimp species, in addition to fat innkeeper worms.

Charlie Endris works on remote-sensing using GIS to analyze habitat change and NERR bio-monitoring pilot studies for Tier 1, emergent vegetation.

Kerstin Wasson monitors oyster recruitment and conducts experiments to determine the status and trajectory of native oyster populations.

John Haskins and Rikke Jeppesen conduct water quality research currently focusing on eutrophication in the slough and are managing the SWMP weather monitoring and the SWMP water quality programs, from which the data are used in conjunction with eutrophication research.

**The following researchers are affiliated with other institutions.**

Aiello, Ivano, Moss Landing Marine Laboratories: examines sediment characteristics relevant to Hester Marsh restoration by collecting sediment cores using a Vibrocore.

Appy, Ralph, Cabrillo Marine Aquarium: collects fish and shrimp to investigate parasites

Anderson, Brian, Siegler, Katie; UC Davis, seine for topsmelt - study pesticide levels in indicator fish.

Baguley, Jeff; University of Nevada: collects mud cores to characterize meifauna esp. harpacticoid copepods

Behesti, Kat; University of California, Santa Cruz: monitors marsh, crabs, and otters with fencing experiments to examine the effects of otters and crabs on salt marsh health.

Bentall, Gena, BeOtterSavy.com: observes otters to characterize the effects of disturbance.

Carlisle, Aron; Dale, Jonathan; Chapple, Taylor: catch small leopard sharks in the Parsons complex, using a gill net, in order to examine physiological response in sharks to different habitat conditions.

Espinosa, Sara; Staedler, Michelle, et al., UC Santa Cruz, USGS: observe otters from land-based sites in order to track sea otter movement and survival (radio-tagged animals) and patterns of distribution, abundance, and behavior (entire Slough population)

Francis, Chris, Stanford University: collects water samples and sediment cores to study the diversity and activity of (de)-nitrifying microbial communities

Hammerstrom, Kamille; Oliver, John; Moss Landing Marine Laboratories: assess effects of Parsons sill on benthic infauna, coffee-can cores and sieving of mud; from shore and boat

Hughes, Brent; UC Santa Cruz: examine effects of otters and crabs on bank erosion and adjacent marshes.

Johnson, Andrew; Monterey Bay Aquarium: tracks sea otter movements; assesses rehabilitated animals; studies otters

Kvitek, Rikk; Spear, Brian; California State University Monterey Bay, use small boat in Parson's complex to map Slough bathymetry to monitor changes resulting from tidal scour

Le, Sam, San Jose State University; uses mapping of Elkhorn Slough bathymetry to monitor changes in the slough channels, resulting from tidal scour

Lidgard, Scott collects bryozoans at Whislestop lagoon to understand gene expression and development in Bugula.

Morrison, Beth, Stanford University: surveys plant transects and collects insect and plant samples to examine how plant-pollinator interactions differ in areas of varying agricultural practices.

Parkin, Jennifer; ESNERR volunteer; monitors salt marsh restoration effects on water birds, and observes the tern colony

Pendleton, Zack; Suraci, Justin; University of California Santa Cruz; studies raccoon use of intertidal habitats

Pien, Cataraina, Moss Landing Marine Laboratories: catches elasmobranchs with gillnets, seines to determine how elasmobranch assemblages have changed over time.

Rahman, Farsana, Drexel University: traps fish with minnow traps to correlate nitrogen isotope levels with nitrate concentration and residence time.

Salkeld, Dan; University of Colorado, Fort Collins; collects ticks to understand patterns of Lyme Disease

Tanner, Karen, Martin Genova; University of California Santa Cruz; transplants and monitors plants to explore the role of mycorrhizae in supporting marsh plants

Tomoleoni, Joseph collects 10-15 invertebrates (jack knife clams, mussels) to compare ease of opening of different types of sea otter prey.

Wise, Maureen (Mo), Moss Landing Marine Laboratories: fly UAV over Estrada/Nurth Marsh, collects sediment and algal samples to understand algal dynamics in Elkhorn Slough.

Wyckoff, Christy, Santa Lucia Conservancy: traps and observes blackbirds to understand movement and habitat use of tri-colored blackbirds.

Yang, Juan; University of California Santa Cruz; collects sediment cores and water samples to understand marsh biogeochemistry

Zabin, Chela, Smithsonian Environmental Research Center: checks oyster restoration experiments to characterize oyster recruitment and physical conditions in Elkhorn Slough.

**Frequent docent researchers**: Shirley Murphy (various bird monitoring programs). Ron Eby (marsh, bird, otter monitoring). Celeste Stanik, Margie Kay, Ken Pollak, Jennifer Parkin (NUT monitoring field and lab work).

**Frequent interns:** Kathleen Hicks (water quality interns).

**II. Physical Structure Descriptors**

**9) Sensor specifications**

ELK NERR deployed EXO2 data loggers at all four sites in 2017. All sondes are the same model with the same probe configuration.

YSI EXO Sonde:

Parameter: Temperature

Units: Celsius (C)

Sensor Type: CT2 probe, Thermistor

Model#: 599870

Range: -5 to 50 C

Accuracy: -5 to 35: +/- 0.01, 35 to 50: +/- .005

Resolution: 0.01 C

Parameter: Conductivity

Units: milli-Siemens per cm (mS/cm)

Sensor Type: CT2 probe, 4-electrode cell with autoranging

Model#: 599870

Range: 0 to 200 mS/cm

Accuracy: 0 to 100: +/- 0.5% of reading or 0.001 mS/cm; 100 to 200: +/- 1% of reading

Resolution: 0.001 mS/cm to 0.1 mS/cm (range dependant)

Parameter: Salinity

Units: practical salinity units (psu)/parts per thousand (ppt)

Sensor Type: CT2 probe, Calculated from conductivity and temperature

Range: 0 to 70 psu

Accuracy: +/- 1.0% of reading pr 0.1 ppt, whichever is greater

Resolution: 0.01 psu

OR

Parameter: Temperature

Units: Celsius (C)

Sensor Type: Wiped probe; Thermistor

Model#: 599827

Range: -5 to 50 C

Accuracy: ±0.2 C

Resolution: 0.001 C

Parameter: Conductivity

Units: milli-Siemens per cm (mS/cm)

Sensor Type: Wiped probe; 4-electrode cell with autoranging

Model#: 599827

Range: 0 to 100 mS/cm

Accuracy: ±1% of the reading or 0.002 mS/cm, whichever is greater

Resolution: 0.0001 to 0.01 mS/cm (range dependent)

Parameter: Salinity

Units: practical salinity units (psu)/parts per thousand (ppt)

Model#: 599827

Sensor Type: Wiped probe; Calculated from conductivity and temperature

Range: 0 to 70 ppt

Accuracy: ±2% of the reading or 0.2 ppt, whichever is greater

Resolution: 0.01 psu

Parameter: Dissolved Oxygen % saturation

Sensor Type: Optical probe w/ mechanical cleaning

Model#: 599100-01

Range: 0 to 500% air saturation

Accuracy: 0-200% air saturation: +/- 1% of the reading or 1% air saturation, whichever is greater 200-500% air saturation: +/- 5% or reading

Resolution: 0.1% air saturation

Parameter: Dissolved Oxygen mg/L (Calculated from % air saturation, temperature, and salinity)

Units: milligrams/Liter (mg/L)

Sensor Type: Optical probe w/ mechanical cleaning

Model#: 599100-01

Range: 0 to 50 mg/L

Accuracy: 0-20 mg/L: +/-0.1 mg/l or 1% of the reading, whichever is greater

20 to 50 mg/L: +/- 5% of the reading

Resolution: 0.01 mg/L

Parameter: Non-vented Level - Shallow (Depth)

Units: feet or meters (ft or m)

Sensor Type: Stainless steel strain gauge

Range: 0 to 33 ft (10 m)

Accuracy: +/- 0.013 ft (0.004 m)

Resolution: 0.001 ft (0.001 m)

Parameter: pH

Units: pH units

Sensor Type: Glass combination electrode

Model#: 599701(guarded) or 599702(wiped)

Range: 0 to 14 units

Accuracy: +/- 0.01 units within +/- 10° of calibration temperature, +/- 0.02 units for entire temperature range

Resolution: 0.01 units

Parameter: Turbidity

Units: formazin nephelometric units (FNU)

Sensor Type: Optical, 90 degree scatter

Model#: 599101-01

Range: 0 to 4000 FNU

Accuracy: 0 to 999 FNU: 0.3 FNU or +/-2% of reading (whichever is greater); 1000 to 4000 FNU +/-5% of reading

Resolution: 0 to 999 FNU: 0.01 FNU, 1000 to 4000 FNU: 0.1 FNU

**LevelQualifier:**

The NERR System-Wide Monitoring Program utilizes YSI data sondes that can be equipped with either vented or non-vented depth/level sensors.  Readings for both vented and non-vented sensors are automatically compensated for water density change due to variations in temperature and salinity; but for all non-vented levelmeasurements, changes in atmospheric pressure between calibrations appear as changes in water depth.  The error is equal to approximately 1.02 cm for every 1 millibar change in atmospheric pressure, and is eliminated for vented sensors because they are vented to the atmosphere throughout the deployment time interval.

Beginning in 2006, NERR SWMP standard calibration protocol calls for all non-vented levelsensors to read 0 meters at a (local) barometric pressure of 1013.25 mb (760 mm/hg).  To achieve this, each site calibrates their levelsensor with a leveloffset number, which is calculated using the actual atmospheric pressure at the time of calibration and the equation provided in the SWMP calibration sheet or digital calibration log.  This offset procedure standardizes each levelcalibration for the entire NERR System.  If accurate atmospheric pressure data are available, non-vented sensor levelmeasurements at any NERR can be corrected.

In 2010, the CDMO began automatically correcting depth/level data for changes in barometric pressure as measured by the Reserve’s associated meteorological station during data ingestion. These corrected depth/level data are reported as cLeveland cLevel, and are assigned QAQC flags and codes based on QAQC protocols. Please see sections 11 and 12 for QAQC flag and code definitions.

**NOTE: older leveldata cannot be corrected without verifying that the leveloffset was in place and whether a vented or non-vented levelsensor was in use. No SWMP data prior to 2006 can be corrected using this method.** The following equation is used for corrected depth/level data provided by the CDMO beginning in 2010:

((1013-BP)\*0.0102)+Depth/Level = cDepth/cLevel.

**Salinity Units Qualifier:**

In 2013, EXO sondes were approved for SWMP use and began to be utilized by Reserves. While the 6600 series sondes report salinity in parts per thousand (ppt) units, the EXO sondes report practical salinity units (psu). These units are essentially the same and for SWMP purposes are understood to be equivalent, however psu is considered the more appropriate designation. Moving forward the NERR System will assign psu salinity units for all data regardless of sonde type.

**Turbidity Qualifier:**

In 2013, EXO sondes were approved for SWMP use and began to be utilized by Reserves. While the 6600 series sondes report turbidity in nephelometric turbidity units (NTU), the EXO sondes use formazin nephelometric units (FNU). These units are essentially the same but indicate a difference in sensor methodology, for SWMP purposes they will be considered equivalent. Moving forward, the NERR System will use FNU/NTU as the designated units for all turbidity data regardless of sonde type. If turbidity units and sensor methodology are of concern, please see the Sensor Specifications portion of the metadata.

**10) Coded variable definitions**

Sampling station: Sampling site code: Station code:

Azevedo Pond AP elkapwq

North Marsh NM elknmwq

South Marsh SM elksmwq

Vierra Mouth VM elkvmwq

**11) QAQC flag definitions**

QAQC flags provide documentation of the data and are applied to individual data points by insertion into the parameter’s associated flag column (header preceded by an F\_). During primary automated QAQC (performed by the CDMO), -5, -4, and -2 flags are applied automatically to indicate data that is missing and above or below sensor range. All remaining data are then flagged 0, passing initial QAQC checks. During secondary and tertiary QAQC 1, -3, and 5 flags may be used to note data as suspect, rejected due to QAQC, or corrected.

-5 Outside High Sensor Range

-4 Outside Low Sensor Range

-3 Data Rejected due to QAQC

-2 Missing Data

-1 Optional SWMP Supported Parameter

0 Data Passed Initial QAQC Checks

1 Suspect Data

2 *Open - reserved for later flag*

3 Calculated data: non-vented depth/level sensor correction for changes in barometric pressure

4 Historical Data: Pre-Auto QAQC

5 Corrected Data

**12) QAQC code definitions**

QAQC codes are used in conjunction with QAQC flags to provide further documentation of the data and are also applied by insertion into the associated flag column. There are three (3) different code categories, general, sensor, and comment. General errors document general problems with the deployment or YSI datasonde, sensor errors are sensor specific, and comment codes are used to further document conditions or a problem with the data. Only one general or sensor error and one comment code can be applied to a particular data point, but some comment codes (marked with an \* below) can be applied to the entire record in the F\_Record column.

General Errors

GIC No instrument deployed due to ice

GIM Instrument malfunction

GIT Instrument recording error; recovered telemetry data

GMC No instrument deployed due to maintenance/calibration

GNF Deployment tube clogged / no flow

GOW Out of water event

GPF Power failure / low battery

GQR Data rejected due to QA/QC checks

GSM See metadata

Corrected Depth/Level Data Codes

GCC Calculated with data that were corrected during QA/QC

GCM Calculated value could not be determined due to missing data

GCR Calculated value could not be determined due to rejected data

GCS Calculated value suspect due to questionable data

GCU Calculated value could not be determined due to unavailable data

Sensor Errors

SBO Blocked optic

SCF Conductivity sensor failure

SCS Chlorophyll spike

SDF Levelport frozen

SDG Suspect due to sensor diagnostics

SDO DO suspect

SDP DO membrane puncture

SIC Incorrect calibration / contaminated standard

SNV Negative value

SOW Sensor out of water

SPC Post calibration out of range

SQR Data rejected due to QAQC checks

SSD Sensor drift

SSM Sensor malfunction

SSR Sensor removed / not deployed

STF Catastrophic temperature sensor failure

STS Turbidity spike

SWM Wiper malfunction / loss

Comments

CAB\* Algal bloom

CAF Acceptable calibration/accuracy error of sensor

CAP Levelsensor in water, affected by atmospheric pressure

CBF Biofouling

CCU Cause unknown

CDA\* DO hypoxia (<3 mg/L)

CDB\* Disturbed bottom

CDF Data appear to fit conditions

CFK\* Fish kill

CIP \* Surface ice present at sample station

CLT\* Low tide

CMC\* In field maintenance/cleaning

CMD\* Mud in probe guard

CND New deployment begins

CRE\* Significant rain event

CSM\* See metadata

CTS Turbidity spike

CVT\* Possible vandalism/tampering

CWD\* Data collected at wrong depth

CWE\* Significant weather event

**13) Post deployment information - Included in annual metadata document**

**Azevedo Pond**

Retrieval DO% DO% pH pH Level Turb Turb Cond Batt

Date of post-cal (0) (100) (7.0) (10.0) (0.0m) (0.0 NTU) (124.0 NTU) (53 ms/cm) (V)

01/03/2017 0.8 103.3 7.17 10.10 1.053 0.00 123.54 58.74 5.4

01/30/2017 0.1 102.8 7.04 9.84 1.116 0.61 126.02 50.45 5.6

02/27/2017 -0.4 101.10 7.12 10.21 1.052 0.92 123.03 50.75 5.6

03/27/2017 0.1 103.0 7.20 10.06 no data 5.75 92.77 47.5 5.7

04/24/2017 0.0 102.8 7.18 10.09 1.105 -0.11 123.26 52.92 5.2

05/22/2017 0.0 101.7 7.17 10.07 1.069 0.43 125.06 51.58 5.3

06/19/2017 0.1 100.6 7.27 10.18 1.021 0.10 123.62 52.75 5.1

07/17/2017 0.4 99.7 7.06 9.96 1.020 0.04 124.58 50.93 5.3

08/14/2017 0.1 99.2 6.69 6.65 0.988 1.39 121.83 51.77 5.5

09/11/2017 0.0 100.5 7.10 9.82 -0.049 0.10 124.80 52.61 5.7

10/10/2017 0.8 99.8 7.02 9.99 nd -0.06 123.9 51.22 5.4

11/06/2017 0.2 102.0 7.16 10.21 1.065 -0.10 123.35 53.50 5.8

12/04/2017 -0.4 99.7 7.21 9.99 1.088 2.57 113.27 52.16 4.1

01/02/2018 1.0 102.4 6.94 9.98 1.094 -0.06 122.75 54.38 6.2

**North Marsh**

Retrieval DO% DO% pH pH Level Turb Turb Cond Batt

Date of post-cal (0) (100) (7.0) (10.0) (0.0m) (0.0 NTU) (124.0 NTU) (53 ms/cm) (V)

01/09/2017 0.3 103.1 7.15 9.98 0.645 0.44 119.34 53.36 out

02/06/2017 0.1 100.9 7.00 9.77 0.526 0.21 122.64 52.77 5.4

03/06/2017 -1.8 102.8 7.06 10.02 0.649 0.15 125.46 51.40 5.2

04/03/2017 -0.9 102.2 7.23 10.15 -0.015 0.23 117.24 51.37 5.5

05/01/2017 -0.7 102.7 7.14 10.06 0.570 -0.03 128.00 51.57 5.2

05/26/2017 28.6 103.4 7.19 10.04 0.562 0.38 121.72 48.09 5.2

06/26/2017 -0.2 102.2 7.20 70.16 0.545 -0.14 125.52 47.84 5.4

07/24/2017 -0.6 100.6 7.17 10.19 0.501 0.27 124.20 51.93 4.8

08/21/2017 -0.6 101.7 7.15 10.12 0.555 -0.11 122.95 44.00 5.7

09/18/2017 0.1 100.4 7.10 10.05 0.540 -0.01 124.50 53.05 6.1

10/16/2017 0.6 101.5 7.13 10.08 0.570 0.04 123.77 52.38 5.6

11/13/2017 0.6 101.7 7.47 10.21 0.609 -0.07 126.13 53.43 5.3

12/08/2017 -0.2 102.6 7.19 10.29 0.656 0.11 122.37 53.05 5.7

01/08/2018 -0.1 100.9 7.14 10.23 0.478 0.03 120.43 53.25 6.2

**South Marsh**

Retrieval DO% DO% pH pH Level Turb Turb Cond Batt

Date of post-cal (0) (100) (7.0) (10.0) (0.0m) (0.0 NTU) (124.0 NTU) (53 ms/cm) (V)

01/25/2017 0.7 102.8 7.07 9.89 -0.681 -0.07 119.94 50.98 5.5

02/23/2017 0.7 101.2 6.97 9.91 -0.899 0.10 118.89 52.25 5.7

03/20/2017 -0.1 100.4 6.98 9.86 -0.926 1.24 112.64 51.64 5.6

04/17/2017 0.1 102.2 6.99 9.59 -0.896 1.35 105.85 52.73 5.4

05/15/2017 -0.9 108.1 7.22 10.12 -0.941 0.05 128.73 52.75 5.1

06/12/2017 0.22 103.8 7.15 10.11 -0.858 0.60 121.85 49.42 5.2

07/10/2017 0.00 96.2 7.09 9.64 -0.952 319.9 401.5 50.04 5.5

08/07/2017 -0.4 100.5 7.09 10.13 -0.943 0.43 124.46 49.80 5.8

09/05/2017 0.7 100.6 7.12 9.88 -0.952 -0.13 124.50 50.97 5.5

10/02/2017 0.2 99.40 7.16 10.04 -0.986 0.16 125.63 53.73 6.1

10/30/2017 -0.2 102.2 7.22 10.14 -0.968 -0.02 122.76 53.22 5.5

11/27/2017 0.4 102.0 7.10 10.05 -0.898 -0.03 123.38 52.89 5.5

12/20/2017 -0.5 103.6 7.34 10.30 -0.817 0.03 125.68 53.03 6.0

01/22/2018 0.4 100.8 7.20 10.07 -0.825 165.40 221.02 52.66 6.0

**Vierra Mouth**

Retrieval DO% DO% pH pH Level Turb Turb Cond Batt

Date of post-cal (0) (100) (7.0) (10.0) (0.0m) (0.0 NTU) (124.0 NTU) (53 ms/cm) (V)

01/17/2017 0.60 101.4 6.96 9.82 -1.592 -0.11 123.95 52.56 6.1

02/13/2017 1.4 105.3 7.15 9.91 -1.578 -0.02 122.30 52.84 4.9

03/09/2017 0.4 103.1 7.19 9.97 -1.572 0.01 126.11 50.38 5.4

04/10/2017 0.4 98.3 7.16 10.30 -1.531 0.15 114.81 53.62 5.2

05/08/2017 0.5 107.4 8.25 11.36 -1.615 -0.06 127.34 51.78 5.5

06/05/2017 0.2 102.2 6.91 9.88 -1.625 0.79 120.17 52.36 5.5

06/29/2017 0.7 no data 7.17 9.91 no data 0.04 125.5 51.82 4.9

07/31/2017 0.8 100.6 6.94 9.95 no data 0.14 122.42 49.33 5.5

08/28/2017 0.2 100.7 7.25 10.12 0.485 0.01 122.69 52.71 5.0

09/25/2017 0.1 99.9 7.34 10.16 -1.660 -0.18 108.22 52.27 5.5

10/23/2017 -0.2 106.5 7.09 10.01 -1.573 -0.02 123.77 53.63 5.2

11/20/2017 0.70 100.50 6.93 9.93 -1.579 0.00 123.52 52.04 5.3

12/18/2017 -0.2 101.9 7.11 9.97 -1.576 0.16 125.25 51.65 5.1

01/15/2018 -0.1 101.5 7.07 9.92 -0.953 0.03 121.88 52.68 5.6

**14) Other remarks/notes**

Data are missing due to equipment or associated specific probes not being deployed, equipment failure, time of maintenance or calibration of equipment, or repair/replacement of a sampling station platform. Any NANs in the dataset stand for “not a number” and are the result of low power, disconnected wires, or out of range readings. If additional information on missing data is needed, contact the Research Coordinator at the reserve submitting the data.

Additional metadata for elkapwq2017:

02/27/2017 10:30:00 to 03/02/2017 12:15:00

11/06/2017 11:45:00 to 11/07/2017 08:30:00

At the following dates/times there was no DO probe deployed on the sonde at the beginning of the deployment. The reserve was short a DO sensor. The sensor was removed from the sonde that was retrieved, brought back to the lab, calibrated and then put on the new sonde. It appears that there are two raw files for this deployment due to this, but they are for the same deployment.

07/17/2017 10:30:00 to 08/14/2017 10:30:00

Level was calibrated wrong. Added 0.047 m to all level records upon retrieval.

Problems with biofouling by hydroids growing in outer tubing cause extreme turbidity spikes this deployment. Spikes rejected and labeled biofouling. pH probe broke on 08/05/2017.

09/11/2017 10:15:00 to 10:45:00

The sonde was deployed at AP on 9/11/17 10:15. The sonde was then accidentally removed when work was being done on the sonde housing during the sampling at 10:30 and 10:45. Data is marked -3 GOW CSM here.

11/06/2017 11:45:00 to 12/04/2017 11:15:00

Although SpCond probe post-cal was reasonable, this cond/sal for this deployment seems shifted down, relative to the preceeding and following deployments.

Additional metadata for elknmwq2017:

01/09/2017 10:45:00 to 11:00:00

The sonde was deployed in the tube at the correct depth but not completely seated correctly. This issue was fixed before the 11:15:00 reading.

03/06/2017 11:15:00 to 04/03/2017 11:30:00

Level off-set was mistakenly left out when calibrating. Hence 0.5382 m was added to correct all level readings for this deployment

Pressbara is consequently also incorrect.

Sonde was out of water on several occasions during this deployment, due to broken tide gates.

04/03/2017 11:45:00 to 05/01/2017 10:30:00

Sonde was out of water on several occasions during this deployment, due to broken tide gates.

05/01/2017 10:30:00 to 05/08/2017 11:00:00

Sonde was deployed next to the permanent station tube, due to leaking tide gates and extremely low water. This data was not collected per current NERR SWMP SOPs. While the other parameters appear to be consistent with station data during this time period due to the well mixed nature of the site, they are not from the correct station location and are not collected per SOPs.

11/13/2017 11:45:00 to 11/14/2017 12:00:00

At the following dates/times there was no DO probe deployed on the sonde at the beginning of the deployment. The reserve was short a DO sensor. The sensor was removed from the sonde that was retrieved, brought back to the lab, calibrated and then put on the new sonde.It appears that there are two raw files for this deployment due to this, but they are for the same deployment.

Additional metadata for elksmwq2017:

01/01/2017 00:00:00 to 01/25/2017 11:30:00

Sonde had accidentally moved and level was no longer correct. On 01/25/2017 John Haskins and Rikke Jeppesen re-leveled the sonde to determine the new level. In order to correct for the sonde movement 0.147 m were subtracted from level in the period.

08/05/2017 01:45:00 to 11/27/2017 13:30:00

Structure holding sonde fell over and sonde was no longer at the exactly correct level. However, the other parameters should still be considered valid.

12/0/2017 09:30:00 to 12/31/2017 23:45:00

Depth was corrected by by making concurrent level readings with a laser level. Measurements were made relative to a pre-leveled reference site. The depth correction was determined by the difference of the readings from the sonde and the laser level. A value of 0.6824 was added to make this correction.

Additional metadata for elkvmwq2017:

03/09/2017 16:45 to 04/10/2017 08:30:00

A rock crab had taken residence in the sonde cage, and most likely caused the extreme turbidity readings for most of this deployment.

04/10/2017 08:45:00 to 04/12/2017 12:30:00

At the following dates/times there was no DO probe deployed on the sonde at the beginning of the deployment. The reserve was short a DO sensor. The sensor was removed from the sonde that was retrieved, brought back to the lab, calibrated and then put on the new sonde.It appears that there are two raw files for this deployment due to this, but they are for the same deployment.

06/29/2017 11:00 to 07/31/2017 09:30

Conductivity sensor broke. Cond. constant on conductivity probe was 5.66 upon retrieval.

06/30/2017 01:15:00 to 07/31/2017 09:30

An organism (crab?) caused the turbidity to have high readings nightly, between 20:00 and 05:00 for about the same amount of time each incident. This is highly unlikely to be increase water turbidity varying with the tides.

07/31/2017 09:45:00 to 08/28/2017 10:15:00

Level was not calibrated before this deployment. In order to correct upon retrieval 2.146m was added to all level readings for this deployment. However, the offset was added after deployment and the sonde was never technically calibrated. This data cannot be assumed to be 100% accurate due to this. Data during this deployment is considered suspect.

09/25/2017 11:45:00 to 10/23/2017 10:00:00

A large rock fish and several crabs were in the sonde cage upon retrieval. These individuals most likely cause the extreme turbidity readings/turbidity spikes for this deployment

10/23/2017 09:45:00 to 11/20/2017 10:00:00

Turbidity spikes in this deployment most likely caused by cancrid crab in sonde cage.

12/18/2017 11:00:00 to 01/15/2018 10:30:00

Level was calibrated incorrectly. 0.628m was subtracted from all level readings.